## Slide 1: Perspectives on Organic Scintillators for Fast Neutron Imaging
Work in Progress Seminar
5 April 2024
Paul Hausladen
(repurposed from NA-22 Workshop on Radiation Detection Materials and Applications)

## Slide 2: Workshop on Radiation Detection Materials and Applications
“The WORDMAp Workshop will involve R&D performers and users from DOE national laboratories, universities, industrial companies, and other organizations engaged in nonproliferation R&D and product development. The purpose is to build a consensus view among subject-matter experts on existing methods for radiation detection, shortcomings in the current state of the art, commercialization issues that are restricting the adoption of new technology, and new materials-related R&D priorities leading to significant improvements in next-generation radiation detectors useful in a variety of nonproliferation applications.” – https://www.bnl.gov/wordmap/
I was present as a ”user” of neutron detector materials rather than a developer

## Slide 3: Detecting Fast Neutrons
For 14.1 MeV neutrons, proton-recoil scintillator is the material of choice (capture reactions do not give practical efficiency)
Can scale to thick, efficient detectors
Direct
Indirect
Scintillator + photosensor
Organometallic semiconductor?
CH2 converter + semiconductor, storage phosphor
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
+
+
+
-
-
-
Neutron
Neutron
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
Neutron
Not demonstrated for neutron detection
Converter limited to proton recoil path length (2-3 mm), eff < 1%

## Slide 4: Integrating Detectors
Observed resolution is width of light distribution (FWHM is 1.53 times height above photosensor) – favors thin detectors
Integrate both neutrons and gamma/x-rays (also produced by neutron sources)
Photodiode has gain ~1 (QE 0.8)
For good images, want more current from neutron exposure than dark current (as much as 1pA/mm2), requires 105 – 106 neutrons/cm2 per second – favors thick detectors for efficiency
h
x
Scintillator
Photosensor
Retroreflector or black

## Slide 5: Efficiency of Proton Recoil Scintillators
For 14.1 MeV neutrons, need ⩾ 3 cm of detector thickness to achieve 10% efficiency
EJ276D: 4.647⨯1022 H atoms/cm3
EJ309: 5.43⨯1022 H atoms/cm3
1/3 of interactions on H, 2/3 of interactions on C

## Slide 6: Resolution vs. Height in Detector
Idealized optical transport with no reflections (or retroreflection) off top surface
PSF

## Slide 7: Resolution vs. Height in Detector
PSF
Height of 15 mm gives 23 mm FWHM

## Slide 8: Resolution vs. Height in Detector
Height of 25 mm gives 41 mm FWHM
PSF

## Slide 9: Resolution vs. Height in Detector
PSF
Height of 35 mm gives 59 mm FWHM

## Slide 10: Resolution vs. Height in Detector
PSF
Height of 45 mm gives 75 mm FWHM

## Slide 11: Contribution to Desired Resolution
For a monolithic scintillator, 90% of light collected on a 1 mm photosensor pixel originates in the lowest 2.8 mm of the column above it.
Most of the light is collected, but in the wrong place. What does this look like?

## Slide 12: Effect of Thickness on Statistics, Resolution
Neutron phantom and idealized measurements by 0.3, 1.0, and 3.0 cm scintillator (5e8 neutron generator at 1 m for 60 s)
3 cm
1 cm
3 mm
No benefit from increased efficiency due to loss of resolution
Mean free paths

## Slide 13: Recovering Resolution in Thick Detectors
Assume pulse-counting, 100 photoelectrons in centroid
Assume 2 mm optical channels
Resolution for thick detectors can be recovered by finding centroid of detected light or channeling light to photosensor

## Slide 14: Pulse-Counting Detectors
When source rates are low enough to distinguish individual interactions in detectors, pulse counting
  Eliminates dark current
  Enables distinguishing neutrons from gamma rays count by count so images do not include x/gamma rays
  Enables centroiding to improve resolution for single interactions from width of light distribution to uncertainty in centroid
  Enables distinguishing the first of multiple interactions through timing
Pulse counting enables the “associated-particle imaging” (API) technique
  Distinguishes transmission neutrons from induced reactions by timing and detected position

## Slide 15: Imaging Using API Neutrons
Neutron source
Neutron Detector
Storage casting in shielded drum
a
n
d+
Make neutrons via reaction
Use coincidence methods to achieve
  excellent transmission contrast with modest source intensity and no physical collimators
  Material information through induced reactions (inelastic gammas/neutrons, induced fission, small angle scattering, kinematic elastic scattering)

## Slide 16: Imaging Using API Neutrons
Neutron source
Neutron Detector
Storage casting in shielded drum
a
n
d+
Time of Flight (ns)
Counts per alpha
Source and object gamma rays
Fission, inelastic scattering, hydrogen elastic neutrons
Transmission and small angle elastic scattering
Coincidence in space

## Slide 17: Example Images
Coincidence techniques enable excellent contrast
Use of late counts and initial trajectory enable labeling of (blue) hydrogen elastic scattering and (red) induced neutron doubles (fission) to assist in identification of materials
No alpha tag
Any alpha tag
18 ns window
Alpha tag via pixels
3 ns window
Overlay of (blue) H-scatter and (red) fission doubles
Maximum attenuation level

## Slide 18: Structures that Guide Light: Precision Arrays
EJ-276D scintillator bonded with 3M ESR reflector (6-mm pixels)
Substantial effort (and material loss) to fabricate
Resolution determined by pixel; pixel size limited by
  Breakage of pixel rows during fabrication
  Light collection for high aspect ratio pixels
  Dead area when fraction of reflector/bonds becomes large
To collect 70% of ideal
10:1 aspect ratio

## Slide 19: Smallest Optical Channels – Proton Track Lengths
Optical channels comparable to the recoil proton track will mean a sizable fraction of protons lose energy in inactive reflector/bond
For a pencil beam of neutrons, recoil proton can go as far as 1 mm in transverse direction
For this approach, 2-3 mm is a reasonable pixel dimension
1 mm

## Slide 20: Structures that Guide Light: Nanoguide
Resolution possible to ~6-10 𝜇m – with this resolution, it is possible to see the transverse projection of proton recoil tracks, can be all scintillator
With this resolution, can in principle,
  Reconstruct for the vertex of a proton track to achieve resolution of tens of microns
  Check that the kinematics are consistent with origination in the neutron generator
Scintillating Nanoguide
Figure provided by Alex Long, LANL

## Slide 21: Scattering in Detector
With 5-cm-thick detector, 70% of hits are recorded away from initial incidence (primarily due to scattering off C)
Fast timing can eliminate most of the additional hits

## Slide 22: Parallax
For modest intensity neutron sources, expect source to be close (1-2 m) to detector
Non-normal incidence will blur resolution in detector, but depth resolution can recover resolution
  For imaging, want single-sided readout
Spread in light may be used to identify depth of interaction

## Slide 23: Review What We Want
For neutron imaging with modest sources, want thick, efficient detectors (~10% efficiency for 14 MeV transmission requires 3 – 5 cm)
  Optical structuring essential to maintaining mm scale spatial resolution in a way that maintains active volume and light collection (small differences in reflectivity are big when reflecting 25 times)
Thick detectors involve additional challenges
  Fast timing (~200 ps FWHM) can identify first of multiple interactions
  Want depth resolution approximately 2.5 times the pixel width to mitigate parallax, ideally from single-sided readout
Fast timing (~ 200 ps FWHM) also desirable to reject random coincidences for API imaging
Do all the above while maintaining neutron-gamma discrimination