## Slide 1: Timepix and Medipix: hybrid pixel detectors which can be applied to multiple applications
M. Campbell1, J. Alozy, R. Ballabriga, R. Bolzonella, E.H.M. Heijne2, X. Llopart, M. Saphorster and L. Tlustos2
CERN, PH Department
1211 Geneva 23
Switzerland
1 Honorary Professor at Glasgow University
2 also with IEAP, Prague
Notes: I am going to talk about the hybrid pixel detector readout chips that have been designed at CERN in the framework of the Medipix2 and 3 collaborations.

## Slide 2: Hybrid Silicon Pixel Detectors
Noise-hit free images possible (high ratio of threshold/noise)
Standard CMOS can be used (follow industry)
Sensor material can be changed (Si, GaAs, CdTe..)
Semiconductor sensor can be replaced by a gas gain grid or MCP
Notes: Mention signal generation Add X section?

## Slide 3: Outline
Origins in High Energy Physics
Introduction to the Medipix and Timepix family
Medipix3 – enabling spectroscopic X-ray imaging
Medipix4 status
Timepix
  Schools
  Dosimetry for manned space flight
Timepix3
  Single layer particle tracking
  Velocity map imaging
  Visible photon imaging and quantum applications
Timepix4
  4DPhoton
Lessons learnt, summary and Conclusion

## Slide 4: Early work on hybrid pixel detectors for HEP
LAA project at CERN led to the development of the first pulse processing pixel prototypes. These early prototypes were developed together with Enz, Krummenacher and Vittoz at EPFL, Lausanne
RD-19 developed hybrid pixel readout for other HEP applications (WA97, NA57)
Delphi pixel detector spun out of RD-19 work
Alice pixel detector was first large scale readout chip in 250nm CMOS using radiation tolerant design techniques
ATLAS and CMS designed their own readout chips based on the 250nm process for the first LHC runs

## Slide 5: One half plane of WA97/RD19 Telescope
36 000 pixels
Omega2 chip
6 ladders of 6 chips
Each chip has 1000 pixels
2 arrays make up one logical plane
[E. Heijne, E. Chesi]
Work carried out by RD19 for WA97.
Basic development funded by CERN-LAA programme.

## Slide 6: A Clean Heavy Ion Physics Event
CERN Experiment WA97 (1995)
5 x 5 cm2 area
7 logical planes
Pixel dimensions 75 x 500 mm2
But…
Trigger precision 1 msec
Radiation tolerance ~ 30kRad

## Slide 7: ATLAS Pixel detector – LHC Run 1
3 barrel layers at ~5, 9 and 12 cm from IP
3 disks on each side 9 ≤ r ≤ 15 cm
Same module layout used everywhere
Module size: 62.4 mm long X 21.4 mm wide
Module read out by 16 FEI3 chips, each serving an array of 18 by 160 pixels
An area of ~1.7 m2 for ~80M pixels

## Slide 8: FE-I3 Module

## Slide 9: Module assembly for ATLAS pixel barrel

## Slide 10: ATLAS high pile-up event – one 25ns ‘frame’
~1cm
~5.5 cm
About 1015 events had been accumulated before the Higgs boson discovery was announced, and only some 1000’s were identified as Higgs boson decays

## Slide 11: Hybrid pixel detectors
Developed initially for first run of the LHC
3 large scale vertex detector systems (ALICE, ATLAS, CMS) operated smoothly in Runs 1 and 2
One large RICH detector system (based on hybrid pixels in a photodetector tube) contributed to LHCb physics in Runs 1 and 2
VELOpix tracker is installed and running in LHCb
Upgrades for ATLAS and CMS (LS3) are in preparation for HL-LHC
In the Medipix2, 3 and 4 Collaborations we have taken the technology into many new fields
This talk will now focus on the readout chips and describes some applications

## Slide 12: The Medipix Collaborations
3 existing Collaborations, hosted at CERN, and comprising only publicly-funded academic and research organisations
Specifications and aims are driven by scientific requirements
ASICs are very expensive both to design and to prototype:
  A single set of masks in a 28nm CMOS process costs ~ EUR 1.5M.
  Highly skilled and specialised manpower is needed for design and verification to avoid multiple submissions
When ASICs are tested and ready for production commercial licenses are negotiated with companies
Much of the dissemination of the technology to other fields happens through the activities of the licensees
2 families of readout chip:
  Medipix – aimed at spectroscopic X-ray imaging
  Timepix – aimed at single particle detection

## Slide 13
Medipix2 (1999 -> ) | Medipix3 (2005 -> ) | Medipix4 (2016 -> )
Albert-Ludwig Universität Freiburg, Germany | Albert-Ludwig Universität Freiburg, Germany | CEA, Paris, France
CEA, Paris, France | AMOLF, Amsterdam, The Netherlands | CERN, Geneva, Switzerland
CERN, Geneva, Switzerland | Brazilian Light Source, Campinas, Brazil | DESY-Hamburg, Germany
Czech Academy of Sciences, Prague, Czechia | CEA, Paris, France | Diamond Light Source, England, UK
ESRF, Grenoble, France | CERN, Geneva, Switzerland | IEAP, Czech Technical University, Prague, Czeciah
IEAP, Czech Technical University, Prague, Czech Republic | DESY-Hamburg, Germany | IFAE, Barcelona, Spain
IFAE, Barcelona, Spain | Diamond Light Source, England, UK | JINR, Dubna, Russian Federation
Mid Sweden University, Sundsvall, Sweden | ESRF, Grenoble, France | NIKHEF, Amsterdam, The Netherlands
MRC-LMB Cambridge, England, UK | IEAP, Czech Technical University, Prague, Czech Republic | University of California, Berkeley, USA
NIKHEF, Amsterdam, The Netherlands | KIT/ANKA, Forschungszentrum Karlsruhe, Germany | University of Canterbury, Christchurch, New Zealand
University of California, Berkeley, USA | Mid Sweden University, Sundsvall, Sweden | University of Geneva, Switzerland
Universität Erlangen-Nurnberg, Erlangen, German | NIKHEF, Amsterdam, The Netherlands | University of Glasgow, Scotland, UK
University of Glasgow, Scotland, UK | Univesridad de los Andes, Bogota, Columbia | University of Houston, USA
University of Houston, USA | University of Bonn, Germany | University of Maastricht, The Netherlands
University and INFN Section of Cagliari, Italy | University of California, Berkeley, USA | University of Oxford, England, UK
University and INFN Section of Pisa, Italy | University of Canterbury, Christchurch, New Zealand | INFN, Italy
University and INFN Section of Napoli, Italy | Universität Erlangen-Nurnberg, Erlangen, German | Chinese Spallation Neutron Source, Dongguan City,China
 | University of Glasgow, Scotland, UK | Brazilian Light Source, Campinas, Brazil
 | University of Houston, USA | Philippines Nuclear Research Institute, Manila, Philippines
 | University of Leiden, The Netherlands | Czech Academy of Sciences, Prague, Czech Republic
 | Technical University of Munich, Germany | Univeristy of Tennessee, Knoxville, USA
 | VTT Information Technology, Espoo, Finland | 

## Slide 14: Commercial Partners
COLLABORATION NAME | Medipix2 |  |  | Medipix3 |  | Medipix4 | 
ASICS | Medipix2 | Timepix | Timepix2 | Medipix3 | Timepix3 | Medipix4 | Timepix4
ADVACAM s.r.o., CZ | X | X | X | X | X |  | X
Amsterdam Scientific Instruments, NL | X | X | X | X | X |  | X
Kromek, UK | X | X | X |  | X |  | 
Malvern-Panalytical, NL | X | X |  | X |  |  | 
MARS Bio Imaging, New Zealand |  |  |  | X |  |  | 
MAX IV, Sweden |  |  |  |  | X |  | 
PITEC, Brazil |  |  | X | X |  |  | X
Quantum Detectors, UK |  |  |  | X |  |  | X
Sequent Logic, USA |  |  |  |  |  |  | X
Sydor Technologies, USA |  |  |  |  |  |  | X
Thermo Fischer Scientific, CZ |  |  |  |  |  |  | X
X-ray Imaging Europe, Germany | X | X |  |  |  |  | 
X-spectrum, Germany |  |  |  | X |  |  | X
X = commercial licensee X = R and D licensee

## Slide 15: ‘Colour’ X-ray imaging

## Slide 16: Cross section of a Hybrid Pixel Detector system (X-ray photon energy deposition)
Sensor dimensions to scale:
55mm pixel pitch
300mm thick sensor
g
Notes: Double counting or missing hits

## Slide 17: Fluorescence in high-Z materials
g
gf
Sensor dimensions to scale:
55mm pixel pitch
300mm thick sensor
Notes: Energy resolution degradation AND missed or double counted hits

## Slide 18
The algorithm for charge reconstruction and hit allocation: Charge Summing Mode
Notes: Note
The analog charge reconstruction (no thresholding before summing).
That as long as the full charge is deposited in a cluster of 2x2 elements it will be correctly reconstructed.

## Slide 19: Medipix3
 | Medipix3RX (2013)
Technology | CMOS 130nm
Pixel Size | 55/110 µm
Sensor Matrix Size | 256 x 256 / 128 x 128
Tile-ability | 3-side
Thresholds | 1/4 Continuous RW
Readout scheme | Sequential RW & Continuous RW
Count Rate (10% deadtime loss) | 4.3 x 106 ph/mm2/s (CSM 110)
Dynamic Range | 25 Ke-
CSM energy resolution (FWHM) | ~ 4.4 KeV (CSM 110)
Readout bandwidth | <1.6 Gbps (8x LVDS)
Power | <1W @1.5V

## Slide 20
Energy Response Function (CdTe 110mm/2mm)
~4.4keV FWHM
Slide courtesy of T. Koenig, KIT
Charge Summing in Spectroscopic X-Ray Detectors With High-Z Sensors; TNS Dec. 2013 4713-4718
Energy (keV)
Normalised differential counts
241Am

## Slide 21: First living human scan…
Phil Butler, CEO of MARS Bio Imaging

## Slide 22: Movie Slice through Phil’s ankle

## Slide 23: MARS wrist scanner
Clinical trials underway at the Hospital for SS, New York

## Slide 24: MARS scan of diseased carotid artery

## Slide 25
Workshop public summary: https://indico.cern.ch/event/1391821/
All large medical equipment suppliers represented: Canon, GE, Philips, Siemens
Also, major research institutes present :Johns Hopkins, Massachusetts General Hospital, Mayo Clinic, Royal Marsden, TU Munich etc.
The Medipix Collaboration plays a ‘pathfinding’ role in this community

## Slide 26
Specxray workshop 2022
Steffen Kappler,
Siemens Healthineers
Member of the Scientific Advisory Committee, Specxray Workshop

## Slide 27: Medipix3  Medipix4
 | Medipix3RX (2013) | Medipix4 (2022)
Technology | CMOS 130nm | CMOS 130nm
Pixel Size | 55/110 µm | 75/150 µm
Matrix Size | 256 x 256 / 128 x 128 | 320 x 320 / 160 x 160
Tile-ability | 3-side | 4-side buttable
Thresholds | 1/4 Continuous RW | 2/8 Continuous RW
Readout scheme | Sequential RW & Continuous RW | 
Count Rate (10% deadtime loss) | 4.3 x 106 ph/mm2/s (CSM 110) | 19 x 106 ph/mm2/s (CSM 140)
Dynamic Range | 25 Ke- | 32 Ke-
CSM energy resolution (FWHM) | ~ 4.4 KeV (CSM 110) | ~ 2.5KeV (CSM 150)
Readout bandwidth | <1.6 Gbps (8x LVDS) | 1,2 Serializers @5.12 or 2.56 Gbps
Power | <1W @1.5V | <1W/cm2 @1.2V

## Slide 28: Status of Medipix4
First assemblies of corrected design receive in July 2024
  75mm sensor pitch
  Testing underway (partly at HIP)
New batch of assemblies received in Feb 2025
  150mm sensor pitch
FPM_SPM 24-bit Mode (Flat field corrected and log scale) 1.0s acquisition @40 KV 10mA tube settings

## Slide 29: Timepix readout chips - single particle detection
 | Timepix | Timepix2 | Timepix3 | Timepix4
Tech. node (nm) | 250 | 130 | 130 | 65
Year | 2005 | 2018 | 2014 | 2019
Pixel size (mm) | 55 | 55 | 55 | 55
# pixels (x x y) | 256 x 256 | 256 x 256 | 256 x 256 | 448 x 512
Time bin (resolution) | 10ns | 10ns | 1.5ns | 200ps
Readout architecture | Frame based (sequential R/W) | Frame based (sequential or continuous R/W) | Data driven or Frame based (sequential R/W) | Data driven or Frame-base (sequential or continuous R/W)
Number of sides for tiling | 3 | 3 | 3 | 4

## Slide 30: Timepix miniaturised readout
IEAP/CTU, Prague

## Slide 31
Advacam s.r.o., Prague
Timepix for schools

## Slide 32: Timepix Demo

## Slide 33: Timepix chip – 60s exposures
Near sea level
34 000 feet

## Slide 34
Several successful pilot projects in the UK, Spain and other countries, have demonstrated that TIMEPIX stimulates and enhances the interest of high school students (especially girls) in STEM subjects
It also improves the motivation of science teachers.
Capturing the interest of high school students in STEM is vital for future recruitment for the microelectronics industry
The CERN & Society Foundation is raising funds to make the technology more widely available in schools
TIMEPIX@school project
https://cernandsocietyfoundation.cern/projects/timepix

## Slide 35
Image of the astronaut Chris Cassidy working near the Timepix USB on the International Space Station (Courtesy of NASA, photo ref. no. iss036e006175)

## Slide 36: Timepix - 4s exposures
South China Sea
South Atlantic Anomaly
University of Houston, IEAP Prague, NASA

## Slide 37: REM Dose Rate Data (mG/min)
University of Houston, IEAP Prague, NASA
Timepix dose rates measured in 2014 on ISS

## Slide 38
Artemis 1 launch Wednesday Nov 16 2022
Carrying 4 Timepix detectors from the Medipix2 collaboration at CERN on board to measure radiation
Part of a larger program at NASA using Timepix based instruments for radiation measurement
3 devices part of HERA radiation dose monitoring hardware. Successful operation from just after launch to just before splashdown
4th device no board Biosentinel cubesat (now at > 70Mkm from earth) Still workign after 1000 days in space!
Timepix on Artemis 1

## Slide 39
Name | Date Flown | Mission | Location | Objective | Vehicle | Number TPX
REM | 2012 | ISS | LEO | Demo | ISS | 5
BIRD | 2014 | Orion EFT-1 | LEO/MEO | Demo/Science | Orion | 2
REM2 | 2018 | ISS | LEO | Ops | ISS | 7
MPT | 2017 | ISS | LEO | Science | ISS | 2
Biosentinel | 2020 | ISS | LEO | Science | ISS | 1
ISS-HERA | 2018 | ISS | LEO | Demo | ISS | 3
AHOSS | 2020 | ISS | LEO | Demo/Ops | ISS | 3
LETS(1) | 2023 | Astrobotic 1 | Lunar Surface | Science | Peregrine | 1
LETS(2) | 2024/5 | Berensheet 2* | Lunar Surface | Science | Berensheet 2 | 1
HERA | 2022 | Artemis 1 | Lunar Orbit | Ops | Orion | 3
Biosentinel | 2022 | Artemis 1 | Solar Orbit | Science | Cubesat | 1
HERA | 2023 | Polaris Dawn | MEO | Science | Crew Dragon | 1
HERA | 2024 | Artemis 2 | Lunar Orbit | Ops | Orion | 6
HERA | 2025 | Artemis 3 | Lunar Orbit | Ops | Orion | 6
ARES | 2025 | Artemis 3 | Lunar Surface | Ops | Starship | >=1
LEIA | ~~2024 | CLPS Lander | Lunar Surface | Science | TBS Lander | 1
ARES | 2026 | Artemis | Lunar Orbit | Ops | Lunar Gateway | 2
*Evaluating mission possibility
7 missions flown, 4 missions next six months, 6 missions manifested, > 23 Timepix in Space to date
Timepix-based flight hardware
S. George (NASA) 10 Years of Timepix in Space – How CERN Detectors are supporting Human Spaceflight https://indico.cern.ch/event/1218130/

## Slide 40: Timepix readout chips - single particle detection
 | Timepix | Timepix2 | Timepix3 | Timepix4
Tech. node (nm) | 250 | 130 | 130 | 65
Year | 2005 | 2018 | 2014 | 2019
Pixel size (mm) | 55 | 55 | 55 | 55
# pixels (x x y) | 256 x 256 | 256 x 256 | 256 x 256 | 448 x 512
Time bin (resolution) | 10ns | 10ns | 1.5ns | 200ps
Readout architecture | Frame based (sequential R/W) | Frame based (sequential or continuous R/W) | Data driven or Frame based (sequential R/W) | Data driven or Frame-base (sequential or continuous R/W)
Number of sides for tiling | 3 | 3 | 3 | 4

## Slide 41: Timepix3 Specs
CMOS node | 130nm
Pixel Array | 256 x 256
Pixel pitch | 55mm
Charge collection | e-, h+
Pixel functionality | TOT (Energy) and TOA (Arrival time)
Preamp Gain | ~47mV/ke-
ENC | ~60e-
FE Linearity | Up to 12ke-
TOT linearity (resolution) | Up to 200ke- (<5%)
TOA resolution* | Up to 1.6ns
Time-walk | <20ns
Minimum detectable charge | ~500e-  2 KeV (Si Sensor)
Power power (1.5V) | 700 mW/cm2
Maximum hit rate | 80Mhits/sec (in data driven)
Readout | Data driven (44-bits/hit @ 5Gbps)
* Thanks to V. Gromov, et al. Nikhef, C. Brezina et al., Bonn

## Slide 42: Timepix3 readout - Trigger-less and event driven
Achievable count rate:
  uniformly distributed events → ~43 Mhits/s/cm2 @5.12Gbps
Full matrix readout: ~800 µs @5.12Gbps
Acquisition time
Address[16-bit]
0xB
Data[28-bits]
Data Packet (48 bits)
48bit
48bit
48bit
48bit
48bit
48bit
End of Command (48 bits)
Shutter
Qin
DataOut
ChipID [32b]
0x71
0xB0

## Slide 43: Pixel Operation in TOA & TOT
Tpeak < 25ns
TOT (10 bits) =4
Preamp Out
Disc Out
Clk (40MHz)
FTOA (4 bits)=7
VCO Clk (640MHz)
TOT Clk (40MHz)
TOA (14-bit)
16383
X
TOA (14 bits)=16383
16384
0
2
3
4
16383
16382
1
Pixel Readout Starts (475ns→ 19 clock cycles)
Global TOA (14-bit)

## Slide 44: Timepix3 miniaturised readout
Advacam s.r.o., Prague

## Slide 45: 3D rendering of traversing particle with delta electron
45 deg
CdTe sensor
2mm thick
Vbias = 130V
Colour (and diameter) indicate charge
Slide courtesy of B. Bergmann, S. Pospisil, IEAP, CTU, Prague

## Slide 46: Setup TPIX3CAM test
A TimePix3 camera was mounted on the ARIADNE prototype TPC we have in Liverpool.
The TPC was filled with 100mb CF4 and the detection/operation principle is the same like in ARIADNE. The light detection efficiency has been directly compared to the EMCCD camera and found to be very similar.
32 cm x 32 cm area read out by a single TPIX3Cam
Slide courtesy of K. Mavrokoridis
Scintillation light

## Slide 47: A selection of cosmic muon events ARIADNE TPIX3Cam
antiproton?
Slide courtesy of K. Mavrokoridis

## Slide 48: An in-vivo treatment monitoring system for ion-beam radiotherapy based on 28 Timepix3 detectors.
This and next 2 slides:
Kelleter, L., Marek, L., Echner, G. et al. An in-vivo treatment monitoring system for ion-beam radiotherapy based on 28 Timepix3 detectors. Sci Rep 14, 15452 (2024). https://doi.org/10.1038/s41598-024-66266-9

## Slide 49: Close up of phantom

## Slide 50: Reconstruction of secondaries
Status: clinical trials underway on a cohort of 50 patients

## Slide 51: Velocity Map Imaging - Timepix3CAM

## Slide 52: Imaging Time of Flight Mass Spectrometry - setup
DOI: (10.1021/acs.analchem.2c04480)

## Slide 53: Imaging Time of Flight Mass Spectrometry - results
DOI: (10.1021/acs.analchem.2c04480)

## Slide 54: Timepix readout chips - single particle detection
 | Timepix | Timepix2 | Timepix3 | Timepix4
Tech. node (nm) | 250 | 130 | 130 | 65
Year | 2005 | 2018 | 2014 | 2019
Pixel size (mm) | 55 | 55 | 55 | 55
# pixels (x x y) | 256 x 256 | 256 x 256 | 256 x 256 | 448 x 512
Time bin (resolution) | 10ns | 10ns | 1.5ns | 200ps
Readout architecture | Frame based (sequential R/W) | Frame based (sequential or continuous R/W) | Data driven or Frame based (sequential R/W) | Data driven or Frame-base (sequential or continuous R/W)
Number of sides for tiling | 3 | 3 | 3 | 4

## Slide 55: Tiling larger areas
Wire Bonding
Target to build large area detectors by combining smaller modules.
The through-silicon vias (TSVs) is the key technology for this paradigm shift.
Medipix4 and Timepix4 use on-chip interposer for bump to pixel redistribution layer (RDL).
The detector is fully sensitive: even above the peripheral circuit and I/O pads.
ASIC
PCB
Sensor
FC
FC
BGA

## Slide 56: Timepix3  Timepix4
 |  |  | Timepix3 (2013) | Timepix4 (2018/19)
Technology |  |  | 130nm – 8 metal | 65nm – 10 metal
Pixel Size |  |  | 55 x 55 µm | 55 x 55 µm
Pixel arrangement |  |  | 3-side buttable 256 x 256 | 4-side buttable 512 x 448
Sensitive area |  |  | 1.98 cm2 | 6.94 cm2
Readout Modes | Data driven (Tracking) | Mode | TOT and TOA | 
 |  | Event Packet | 48-bit | 64-bit
 |  | Max rate | <0.43 Mhits/mm2/s | <3.58 Mhits/mm2/s
 |  | Max pix rate | 1.3kHz/pixel | 10.6kHz/pixel
 | Frame based (Imaging) | Mode | PC (10-bit) and iTOT (14-bit) | CRW: PC (8 or 16-bit)
 |  | Frame | Zero-suppressed (with pixel addr) | Full Frame (without pixel addr) CRW (8-bit / 16-bit) Up to 44 KHz frame @8b
 |  | Max count rate | 82 Ghits/cm2/s | ~800 Ghits/cm2/s
TOT energy resolution |  |  | < 2KeV | < 1Kev
Time resolution (bin size) |  |  | 1.56ns | ~200ps
Readout bandwidth |  |  | ≤5.12Gb (8 x SLVS@640 Mbps) | ≤163 Gbps (16 x 10.24 Gbps)
Target global minimum threshold |  |  | <500 e- | <500 e-

## Slide 57: Timepix4 Pixel Schematic
Global threshold
Front-end
Leakage
Current
compensation
Preamp
Input
pad
1 pixel
VCO
@640MHz
Super pixel
Control
voltage
8 pixels
640MHz
TpA
TpB
TestBit
MaskBit
Front-end
Counters
&
Latches
Time stamp
14-bits
Synchronizer
&
Clock gating
OP Mode
5-bit Local
Threshold
~50mV/ke-
2.5fF
3fF
Sync
Readout
Data out
to EOC
8-bits
clock
(40MHz)
8-bits
4
8-bits
ADB
32 pixels
8
Clock
local
 | bits
ADDR | 9
ToA | 16
ufToA_r | 4
ufToA_f | 4
fToA_r | 5
fToA_f | 5
ToT | 12
Pileup | 1
SPG

## Slide 58: Pixel Operation in TOA & TOT [DD]
Preamp Out
Disc Out
Clk (40MHz)
VCO Clk (640MHz)
TOT Clk (40MHz)
TOA (16-bit)
Global TOA (16-bit)
UFTOA_Start (4-bits)
UFTOA_Stop (4-bits)
65534
X
TOA (16 bits)=65534
Tpeak < 25ns
TOT (10 bits) =4
65535
0
2
3
4
65534
65533
1
FTOA_R (4 bits)=7
FTOA_F (4 bits)=11
0
X
8
X

## Slide 59: Timepix4 Readout Modes : Data-Driven
Zero-suppressed continuous data-driven
  Output bandwidth from 40 Mbps (2.6 Hz/pixel) to 160 Gbps (10.8 KHz/pixel)
  Uses Aurora 64b/66b standard encoding communication protocol
4 External Sync Inputs to synchronize/align external signals with data
Acquisition time
64bit
64bit
64bit
64bit
64bit
64bit
Shutter
External Sync Inputs
DataOut
Address: 18 bits
Time: 29 bits
Energy: 21 bits
Qin (FE)
64bit
64bit

## Slide 60: Timepix4 Floorplan
29960 µm
TSVs
TSVs
24700 µm
q
28160 µm (512 x 55 µm)
24640 µm (448 x 55 µm)
TSV_TOP (460 um)
WB_TOP (900 um)
TSV_MID (920 um)
TSV_BOT (460 um)
WB_BOT (900 um)
13158.4 µm
(256 x 51.4 µm)

## Slide 61: Timepix4 Bottom left detail
BOTTOM Periphery
BOTTOM Matrix
BOTTOM
WB Extender
Dicing lane #1
Dicing lane #2
460 µm

## Slide 62: Timepix4 assembly (300mm Si sensor)

## Slide 63: Timepix4 – works! 
10 s exp. 90Sr
Threshold ~ 800e-
6.1 M packets @ 5 Gbps
5 ms timeslice

## Slide 64: Photon counting image Timepix4
Timepix4 paper:
X. Llopart et al 2022 JINST 17 C01044

## Slide 65: 4Dphoton: Integrate Timepix4 in a photo tube
Concept already proven with 4 Timepix chips
See: J Vallerga et al. https://iopscience.iop.org/article/10.1088/1748-0221/9/05/C05055
Ongoing effort with Timepix4 started
See: M. Fiorini et al. https://iopscience.iop.org/article/10.1088/1748-0221/13/12/C12005/pdf

## Slide 66: Energy calibration with sources
https://indico.global/event/8935/contributions/85485/attachments/39904/74219/iWoRiD_2024_Bolzonella.pdf

## Slide 67: Calibration of on pixel VCOs
https://indico.global/event/8935/contributions/85485/attachments/39904/74219/iWoRiD_2024_Bolzonella.pdf

## Slide 68: Integration into photo-sensitive tube
https://indico.global/event/8935/contributions/85485/attachments/39904/74219/iWoRiD_2024_Bolzonella.pdf

## Slide 69: TSV processing of Timepix4 - 3 on-going activities
Supplier | IZM (D) | LETI (F) | PTI (TW)
Wafer size (mm) | 200 | 200 | 300
Front side UBM | Yes | No | Yes
Front side bumps | No | No | Yes
TSV diameter/ thickness (um) | 60/120 | 60/120 | 38/180
Rear side RDL | Yes | Yes | Yes
Status | Second batch of wafers just completed | First batch of wafers just completed | Setup wafers almost done

## Slide 70
Examples of other applications
X-ray materials analysis with diffraction
Photon counting cameras at synchrotrons
Time resolved cameras at synchrotrons
X-ray non-destructive testing
X-ray dosimetry - dosepix chip development
Waste sorting using X-rays
Low Energy Electron Microscopy
Transmission electron microscopy
Dose deposition tracking in hadron therapy
Numerous satellite systems dedicated to space weather observation

## Slide 71
Applications for CERN/Physics
  LHCb VELOpix chip is directly derived from Timepix3
  LHCb Timepix3 telescope – 80 Mhits/cm2/sec
  Sensor studies for CLIC/LHCb
  Background radiation monitoring at ATLAS and CMS
  Beam Gas Interaction monitor CERN PS
  Beam monitoring in UA9
  Positron annihilation in Aegis
  ASACUSA experiment
  Breit-Wheeler experiment at RAL
  Beta particle channeling in ISOLDE
  Forward physics using Timepix3?
  Axion search at CAST (with InGrid)
  Large area TPC (with InGrid)
  Transition radiation measurements for ATLAS
  GEMPIX development for radiation therapy beam monitoring
  GEMPIX for 55Fe waste management
  Developments for CLIC: CLICpix, CLICpix2, C3PD

## Slide 72: The Picopix project
PicoPix is a large-scale hybrid pixel tracking particle detector designed in 28nm CMOS:
Large area  3-side buttable with 384 x 256 pixel array
Spatial resolution  square pixel size of 50 μm
Time measurements  <50 psrms per Hybrid (ASIC + Sensor)
Radiation hardness  < 1 Grad with full TMR (Triple Modular Redundancy) logic design
Data reduction  On-chip programmable event clustering, veto, filtering and sorting
Event rate  Up to 96 events/chip/25ns  ≤3.84∙109 events/chip
High speed serializer  Data-driven readout with up-to 4 x 25.6 Gbps fast links
Design Approach:
Built from gained experienced from previous hybrid pixel detectors tracking detectors designed in EP-ESE-ME:
  Timepix (2007)  Timepix3 (2013)  Velopix1 (2016)  Timepix4 (2020)
  CERN and Nikhef design team
Directly profiting from CERN EP-RD program:
  WP5.3: Intelligence on detector, WP5.2: IP blocks, WP5.5: 3D interconnects
  WP6: High Speed links

## Slide 73: Potential Applications of PicoPix
LHCb VELO Upgrade 2 & Fast-Sensor R&D
  High-precision tracking for High-Luminosity LHC (HL-LHC)
  Timing resolution below 50 ps to distinguish multiple collisions
  Radiation-hard hybrid silicon sensors for extreme environments
SY-BI Beam Loss Monitoring
  Bunch-by-bunch beam loss detection
  Improved diagnostics for LHC and future e⁺e⁻ colliders
Medipix3 Collaboration (Medical & Industrial Applications)
  Hadron therapy beam monitoring in cancer treatment
  X-ray diffraction & electron microscopy for material science
  Single-layer Compton cameras for homeland security and SPECT imaging
  Quantum sensing using entangled particle detection

## Slide 74
Pixel Matrix
384 x 256
(pixel size 50mm x 50mm)
Periphery
SPRINT
19.2 mm
12.8 mm
<3.5mm
Picopix Floorplan
SPRINT = Silicon Photonics Radiation-tolerant Integrated Transmitter

## Slide 75
Front-end: Charge Sensitive Amplifier
Front-end: Quenching element, regeneration circuit
Picopix front-end
Notes: A Charge Sensitive Amplifier (CSA) is an electronic circuit designed to convert an input charge (from a sensor or detector) into a proportional voltage output. It is especially common in radiation detection where detectors deliver small charge in response to incident particles.
It presents a low input impedance and a gain (Volts/Charge) which is determined by the inverse of the value of the feedback capacitor (1/CF).

## Slide 76: Picopix Time Measurement  1 TDC every 4 pixels
  Free running Digital Control Oscillator
    Period 400ps (2.5GHz)
    UFTOA: ~30-40 ps bin derived from state of inner nodes of VCO
  Auto-calibrated TDC at each measurement:
    P, V, T, Rad and On-chip variations on a large-scale chip are critical for this resolution
  Clock40 designed as an improved precision Timing Clock Tree with Digital DLL (~10 ps clock skew)
  Clock320 used to operate the TDC faster. 1 Measurement in <50ns (Timepix4 was ~300ns)

## Slide 77: On pixel event processing ➞ Data reduction
Data filtered already at the level of the pixel-array
  Bandwidth reduction towards the periphery
  Significant power and IR-drop reduction
Further reduction and compression in the periphery
Filtering disable
Filtering large clusters
Filtering small clusters
  Design Status: RTL complete
Clustering and filtering

## Slide 78: Periphery data-path
Full Time calibration:
  DLL clock skew calibration (Row and Column)
  Unique time-walk correction (programable LUT 32x32)
  TOA event calibration (24.4ps bin)
  10b or 6b fine-time output
  Built-in TOA extender up to 32-bits(@40MHz)
  Digital T0Sync shift adjustment (24.4ps bin and 100ns range)
Event Filtering:
  Configurable TOT range
  Configurable TOA range
  Hitmap shape (single pixels, L shape, + shape,…)
On-chip XY decoding
Single packet formatting:
  To Sorter: 24-bit, 36-bit, 72-bit
  To FIFO: 72b_PC_iTOT, 72b_raw and Jumbo
PACKET PROCESSOR

## Slide 79
Lessons learnt
Philippe Starck
Juicy Salif citrus squeezer
Finnish Design Shop
Some compare HEP with a famous Italian design:
Beautiful
Expensive
But almost completely useless
With the Medipix Collaboration we have demonstrated that the pursuit of advanced detector readout concepts can impact many other areas of science
But some elements are required:
Imagination/vision
A willingness to pool ideas and resources
Openness to the requirements of others
Highly skilled engineers
Organisation, tenacity, patience
Fair and workable licensing

## Slide 80: Summary and Conclusions
Hybrid pixel detectors were developed as tracking detectors of LHC
The Medipix and Timepix family of chips span over 20 years. With each new generation more functionality has been added to the ASIC
Spectroscopic X-ray imaging in medicine is becoming a reality
The Timepix3 and 4 ASICs when combined with a Si sensor provide noise-hit-free bubble-chamber-like imaging with always-on triggerless readout.
When combined with an appropriate convertor material, single visible photon hits can be streamed out – extremely useful in quantum sensing applications
Timepix4 (time stamp within 200ps bin) can be tiled on 4 sides using TSVs. Medipix4 to follow.
Picopix improves time stamping and introduces on-pixel and on-chip data filtering
Time, resources and patience ultimately pay off

## Slide 81: Invitation

## Slide 82
Thank you for your attention!
Medipix3RX images: S. Procz et al.
Timepix4 paper:
X. Llopart et al 2022 JINST 17 C01044

## Slide 83
Backup

## Slide 84: Timing
Target resolution of <20psrms track time is required to distinguish PV
Single plane resolution of <50psrms  σ2sensor (40psrms) + σ2ASIC
σ2ASIC (30psrms) σ2analogFE + σ2conversion + σ2clock
  σ2conversion   TDCbin = 40ps  11.5 psrms
  σ2clock  Reference clock at pixel level < 10 psrms
  σ2analogFE  < 25 psrms

## Slide 85: Fundamental limits to time resolution in analog front-end
R.Ballabriga, The Timepix4 analog front-end design: Lessons learnt on fundamental limits to noise and time resolution in highly segmented hybrid pixel detectors https://doi.org/10.1016/j.nima.2022.167489
  , are the front-end input, feedback, and output capacitances, respectively.
  is the input transconductance
  is the threshold of the comparator
→ Minimize the front-end output capacitance.
→ Increase the analog pixel power consumption, but a large power drop over the ASIC worsens the timing performances!
→ The front-end input capacitance must be minimized, which depends on the choice of the sensor, pixel pitch…
→ Maximize the input charge and decrease the threshold.